Molecular Testing for Malaria Standard Operating Procedure (SOP)

DNA Extraction by Chelex

Document ID: EXT-02

In association with the Mekong Molecular Surveillance for Drug Resistant Malaria program, supported by USAID

WWARN Clinical Group

WorldWide Antimalarial Resistance Network (WWARN)

www.wwarn.org

Suggested citation: Clinical Group, WWARN. 2015. Molecular Testing for Malaria Standard Operating Procedure DNA extraction by Chelex v1.1

Procedure ID: EXT-02

This procedure was developed by:

WWARN Clinical Group

WorldWide Antimalarial Resistance Network

University of Maryland School of Medicine

Mahidol University

In association with the Mekong Molecular Surveillance for Drug Resistant Malaria program, supported by USAID

Global Scientific Solutions

Version History

| Version number | Revision(s) & reason for amendment | Date of release |
| --- | --- | --- |
| 1.0 | Creation of procedure | 3 Jan 2012 |
| 1.1 | Updated to v1.1 – corrected procedure errors; format different from other procedures; overly long | 10 Mar 2015 |

For more information, contact: 	
clinical@wwarn.org

WorldWide Antimalarial Resistance Network (WWARN)

www.wwarn.org

### Molecular Testing for Malaria: Overview of Standards

A number of consensus-adopted standards have been used in the development of this document. These include:

- Recommended Genotyping Procedures (RGPs) to identify parasite populations (World Health Organization, 2007).

- MORU Standard Operating Procedure: DNA Extraction (Nakeesathit, Pagomrat, Tanomsing, & Hanchana, 2001).

The currently available texts were developed from work reported in research articles by: (Kyes, Craig, & Marsh, 1993); (Plowe, Diimde, Bouare, & Doumbo, 1995); (Rubio, J., L., Garcia, M., & I., 1999); (Färnert, et al.)

### Signature Page

By signing this page, staff members providing malaria recrudescence versus reinfection testing confirm they have read this SOP and guarantee to implement the procedures contained within.

| Name | Designation/affiliation | Signature | Date of signing |
| --- | --- | --- | --- |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |

### Version History

This SOP may be adapted to suit the particular needs of the individual laboratory. It is important to bear in mind that the purpose of an SOP is to ensure testing quality and result reproducibility.

| Version number | Revision(s) & reason for amendment | Date of approval | Approved by (lab supervisor / manager) |
| --- | --- | --- | --- |
| 1.1 | Corrected errors | 09/02/2015 | Mehul Dhorda |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |

### Contents

Molecular Testing for Malaria: Overview of Standards	3

Signature Page	4

Version History	5

Contents	6

1.	Scope	8

2.	Abbreviations	8

3.	Personnel qualifications	8

3.1.	Medical fitness	8

3.2.	Education and training	8

4.	Procedure	8

4.1.	Principle	8

4.2.	Samples	9

4.3.	Required Materials and Equipment	9

4.4.	Procedural steps	9

5.	Quality Control	11

5.1.	Negative Control	11

6.	Procedure limitations	11

7.	Interpretation and Reporting of Results	11

8.	Safety Precautions	11

9.	Bibliography	12

Appendix A – Bench Top Reference	14

Appendix B – Sample Worksheet	15

# Scope

This SOP describes the method of extracting genomic DNA from Plasmodium spp. by Chelex-100/Boiling.

# Abbreviations

DBS	Dried Blood Spot

ml	milliliter

PBS	Phosphate Buffered Saline

PCR	Polymerase chain reaction

SOP	Standard Operating Procedure

L	microliter

# Personnel qualifications

## Medical fitness

Occupational health programs should be in place to monitor/address staff vaccinations and deal with exposures to potentially infected materials.

## Education and training

Training must be given on the following topics:

- Wearing and use of personal protective equipment and clothing;

- Handling of potentially infectious materials;

- Prevention of incidents and steps to be taken by workers in case incidents (including biological, chemical, electrical and fire hazards) occur;

- Procedures;

- Waste management;

- Impact of results for patient management and research.

Training must be provided:

- When a new staff member takes up post;

- Annually;

- When there is a change in conditions or best practices.

# Procedure

## Principle

Plasmodium DNA from dried blood spots is extracted by heating samples in a suspension of 5 % Chelex 100. Heating to almost 100 ºC in an alkaline suspension disrupts the cell membrane and breaks down proteins including heat-labile enzymes while Chelex 100, an ion chelator, limits destruction of the DNA by inactivating nucleases and chelating heavy metals that may damage parasite DNA.

This method makes it possible to obtain parasite DNA suitable for PCR. It is a fast, cheap, and effective method of DNA extraction. As this is the first step in the PCR process, it is important to maintain sterile technique to prevent contamination of extracted DNA.

## Samples

This method is designed to be performed on dried blood spot (DBS) samples (see SOP-01 DBS Sample Collection.

It is critical to ensure the DBS sample is of the highest quality. Poor quality samples will lead to poor quality results.

## Required Materials and Equipment

1.5 ml microcentrifuge tubes

1-20 l single channel automatic pipettes

100-200 l single channel automatic pipette

1000 l single channel automatic pipette

Filter pipette tips for the above pipettes

Fine tip marker pens

Ball point pen

Paper towels or wipes

Distilled water

Phosphate Buffered Saline (PBS)

Bleach (5 %) in a beaker or wash bottle

Distilled water in a beaker or wash bottle

Ethanol (70 %) in a beaker or wash bottle

Chelex®-100 Resin

Scissors or 1/8 inch hole punch (plus spare filter paper if using a punch)

Timer

Microcentrifuge

Heating block or waterbath at 56 ºC

Waterbath at boiling temperature (96 ºC or above)

Vortex

## Procedural steps

Important points to remember: - Ensure the scissors are thoroughly cleaned before beginning the procedure, in between cutting filter papers and at the end of the procedure. Unclean scissors can lead to cross contamination of samples and poor quality results. - Ensure pipette tips are of a high quality, sterile and endonuclease free. - Do not touch pipette tips. - Make sure pipettes are calibrated and cleaned regularly.

- Print out a PCR worksheet and record the sample ID of each DBS to be tested on a separate numbered line.

- Gather all required supplies. NB. If samples have been stored at +4 ºC or -20ºC they must be brought to room temperature in the sample bag prior to opening.

- Gather all required supplies.

Make Chelex reagent (20 %) by adding distilled water (e.g. 2 g Chelex with 10 ml distilled water). NB. Chelex reagent should be made fresh each day it is required.

- Clean the scissors or punch by dipping in ethanol (70%) and passing through a flame.

- Label an appropriate number of 1.5 ml microcentrifuge tubes (label both the lid and the side of the tube) with the worksheet number and sample ID.

- Cut 2-3 pieces of 3 mm x 3 mm or punch a 3 mm disk (holds approx. 3-5 l of dried blood) from the filter paper and put it into the corresponding 1.5 microcentrifuge tube or well of a 96-well microtitre plate. NB. Clean the scissors between each sample as detailed in step 5. Clean the punch by punching clean filter paper 3 times.

- Add 1 ml PBS. NB. Ensure filter papers are soaked in buffer.

- Incubate at room temperature for 10 min.

- Centrifuge at 14,000 rpm for 2 min and discard supernatant using a clean pipette tip for each sample.

- Add 1 ml PBS

- Centrifuge at 14,000 rpm for 2 min and discard supernatant using a clean pipette tip for each sample.

- Add 150 l of nuclease-free water.

- Add 50 l of 20 % Chelex.

- Incubate at 99 ºC for 10 min.

- Centrifuge at 14,000 rpm for 1 min.

- Store supernatant at +4ºC for use in PCR. NB. If storing samples for longer than a day, transfer supernatant into a fresh microcentrifuge tube and stored at -20 ºC.

# Quality Control

## Negative Control

To test for cross-contamination, each batch of 10 samples should contain at least 1 negative control. The negative control consists of a section of plain filter paper cut in the same way as the DBS samples.

# Procedure limitations

Successful extraction of DNA is dependent upon the quality and quantity of DNA in the DBS sample, quality of laboratory reagents, equipment and supplies, and the implementation of good quality clinical laboratory practice according to this SOP.

# Interpretation and Reporting of Results

The extracted DNA sample is used as template DNA for subsequent PCR reactions. There are no reporting requirements at this stage apart from the rejection of poor quality samples.

# Safety Precautions

The dried blood spot samples used in this procedure should be treated as potentially infectious biology material under BSL2 conditions. Universal precautions for class 2 pathogens apply.

Always practice universal precautions (treat all patient specimens as potentially infectious material):

- Wear good quality, single-use, disposable medical examination gloves.

- Wear a laboratory coat or gown.

- Wash hands after removal of gloves.

Dispose of medical/laboratory waste in the appropriate manner:

- Contaminated waste must be disposed of immediately after use into a proper waste container.

- Spills should be cleaned using 10% bleach.

# Bibliography

Calderaro, A., Piccolo, G., Perandin, F., Gorrini, C., Peruzzi, S., Zuelli, C., et al. (2007). Genetic polymorphisms influence Plasmodium ovale PCR detection accuracy. Journal of Clinical Microbiology, 45(5), 1624-1627.

Cattamanchi, A. A., Kyabayinze, D., Hubbard, A., & Rosenthal, P. J. (2003). Distinguishing recrudescence from reinfection in a longitudinal antimalarial druf efficacy study: comparison of results based on genotyping of msp-1, msp-2, and glurp. American Journal of Tropical Medicine, 68(2), 133-139.

Centers for Disease Control and Prevention. (2006). Blood collection - finger prick. Retrieved August 17, 2010, from CDC HIV Training: http://wwwn.cdc.gov/dls/ila/hivtraining/trainersguide/pdf/presentations/Module8Presentation.pdf

Clinical and Laboratory Standards Institute. (2007). Blood collection on filter paper for newborn screening programs; approved standard - fifth edition. CLSI document LA4-A5, 27(20).

Falk, N., Maire, N., Sama, W., Owusu-Agyei, S., Smith, T., & Beck, H.-P. a. (2006). Comparison of PCR-RFLP and Genescan-based genotyping for analyzing infection dynamics of Plasmodium falciparum. American Journal of Tropical Medicine., 74(6), 944-950.

Färnert, A., Arez, A. P., Babiker, H. A., Beck, H. P., Benito, A., Björkman, A., et al. (n.d.).

Färnert, A., Arez, A. P., Babiker, H. A., Beck, H. P., Benito, A., Björkman, A., et al. (n.d.). Genotyping of Plasmodium falciparum infections by PCR: a comparative multicentre study. Trans R Soc Trop Med Hyg., 95.

Kyes, S., Craig, A. G., & Marsh, K. a. (1993). Plasmodium falciparum: a method for the amplification of S antigens and its application to laboratory and field samples. Experimental Parasitology, 71, 473-483.

Nakeesathit, S., Pagomrat, W., Tanomsing, N., & Hanchana, S. a. (2001). DNA Extraction. Bangkok: Mahidol Oxford Tropical Medicine Research Unit.

Plowe, C. V., Djimde, A., Bouare, M., & Doumbo, O. a. (1995). Pyrimethamine and proguanil resistance-conferring mutations in Plasmodium falciparum dihydrofolate reductase: polymerase chain reaction methods for surveillance in Africa. American Journal of Tropical Medicine and Hygiene., 52, 565-568.

Program for Appropriate Technology in Health (PATH). (2005). RBP-EIA: collecting, processing, and handling venous, capillary, and blood spot samples. Seattle: PATH.

QIAGEN. (2007). QIAamp DNA mini and blood mini handbook Second edition. QIAGEN.

Rubio, J. M., J., R. P., L., B. M., Garcia, M., M., M., & I., E. M. (1999). Semi-nested. multiplex polymerase chain reaction for detection of human malaria parasites and evidence of Plasmodium vivax infection in Equatorial Guinea. American Journal of Tropical Medicine and Hygiene., 60, 183-187.

U.S. Department of Health and Human Services. (2007). Biosafety in microbiological and biomedical laboratories. 5th. Washington, District of Columbia, USA: U. S. Government Printing Office.

U.S. Department of Health and Human Services. (2007). Biosafety in microbiological and biomedical laboratories (5th ed.). Washington, District of Columbia, USA: U. S. Government Printing Office.

Watcharee, P., Naowarat, T., Supatchara, N., & Sarun, H. a. (2009). Gel Electrophoresis. Bangkok: MORU.

World Health Organization. (2007). Methods and techniques for clinical trials on antimalarial drug efficacy: genotyping to identify parasite populations. Amsterdam: WHO.

World Health Organization. (2007). Recommended Genotyping Procedures (RGPs) to identify parasite populations. Amsterdam: Medicines for Malaria Venture and World Health Organization.

World Health Organization. (2008). Consultation on Technical and Operational Recommendations for Clinical Laboratory Testing Harmonization and Standardization. Geneva: World Health Organization.

Worldwide Antimalarial Resistance Network. (2010). Filter paper preparation v1.0 (SOP ID: MOL03/CLIN06). Worldwide Antimalarial Resistance Network (WWARN).

Zwetyenga, J., Rogier, C., Tall, A., Fontenille, D., Snounou, G., & Trape, J.-F. a.-P. (1998). No influence of age on infectious complexity and allelic distribution in Plasmodium falciparum infections in Ndiop, a Senegalese village with seasonal, mesoendemic malaria. American Journal of Tropical Medicine and Hygiene., 59(5), 726-735.

### Appendix A – Bench Top Reference

DNA Extraction by Chelex - Document ID: BTR-03

- Print out a PCR worksheet and record the sample ID of each DBS to be tested on a separate numbered line.

- Gather all required supplies. NB. If samples have been stored at +4 ºC or -20ºC they must be brought to room temperature in the sample bag prior to opening.

- Gather all required supplies.

Make Chelex reagent (20 %) by adding distilled water (e.g. 2 g Chelex with 10 ml distilled water). NB. Chelex reagent should be made fresh each day it is required.

- Clean the scissors or punch by dipping in ethanol (70%) and passing through a flame.

- Label an appropriate number of 1.5 ml microcentrifuge tubes (label both the lid and the side of the tube) with the worksheet number and sample ID.

- Cut 2-3 pieces of 3 mm x 3 mm or punch a 3 mm disk (holds approx. 3-5 l of dried blood) from the filter paper and put it into the corresponding 1.5 microcentrifuge tube or well of a 96-well microtitre plate. NB. Clean the scissors between each sample as detailed in step 5. Clean the punch by punching clean filter paper 3 times.

- Add 1 ml PBS. NB. Ensure filter papers are soaked in buffer.

- Incubate at room temperature for 10 min.

- Centrifuge at 14,000 rpm for 2 min and discard supernatant using a clean pipette tip for each sample.

- Add 1 ml PBS

- Centrifuge at 14,000 rpm for 2 min and discard supernatant using a clean pipette tip for each sample.

- Add 150 l of nuclease-free water.

- Add 50 l of 20 % Chelex.

- Incubate at 99 ºC for 10 min.

- Centrifuge at 14,000 rpm for 1 min.

- Store supernatant at +4ºC for use in PCR. NB. If storing samples for longer than a day, transfer supernatant into a fresh microcentrifuge tube and stored at -20 ºC.

### Appendix B – Sample Worksheet

| DNA Extraction by Chelex |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Staff Initials |  |  |  | Date |  |  |  |
| Number | Patient ID |  | Sample ID |  | Storage | Comments |  |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |
| Negative Control |  |  |  |  |  |  |  |