Molecular Testing for Malaria Standard Operating Procedure (SOP)

msp1 PCR

Document ID: EXT05

WWARN Clinical Group

WorldWide Antimalarial Resistance Network (WWARN)

www.wwarn.org

In association with the Mekong Molecular Surveillance for Drug Resistant Malaria program, supported by USAID

Suggested citation: Clinical Group, WWARN 2015. msp1 PCR v1.1

Procedure ID: EXT05

This procedure was developed by:

WWARN Clinical Group

WorldWide Antimalarial Resistance Network

University of Maryland School of Medicine

Mahidol University

In association with the Mekong Molecular Surveillance for Drug Resistant Malaria program, supported by USAID

Global Scientific Solutions

Version History

| Version number | Revision(s) & reason for amendment | Date of release |
| --- | --- | --- |
| 1.0 | Creation of procedure | 3 Jan 2012 |
| 1.1 | Minor text changes | 10 Mar 2015 |

For more information, contact: 	
clinical@wwarn.org

WorldWide Antimalarial Resistance Network (WWARN)

www.wwarn.org

### Molecular Testing for Malaria: Overview of Standards

A number of consensus-adopted standards have been used in the development of this document. These include:

- Recommended Genotyping Procedures (RGPs) to identify parasite populations (World Health Organization, 2007).

- MMV Guide to Malaria Genotyping (World Health Organization, 2008).

- Biosafety in Microbiological and Biomedical Laboratories (BMBL) 5th Edition (U.S. Department of Health and Human Services, 2007).

### Signature Page

By signing this page, staff members providing malaria recrudescence versus reinfection testing confirm they have read this SOP and guarantee to implement the procedures contained within.

| Name | Designation/affiliation | Signature | Date of signing |
| --- | --- | --- | --- |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |
|  |  |  | [DD/MM/YYY] |

### Version History

This SOP may be adapted to suit the particular needs of the individual laboratory. It is important to bear in mind that the purpose of an SOP is to ensure testing quality and result reproducibility.

| Version number | Revision(s) & reason for amendment | Date of approval | Approved by (lab supervisor / manager) |
| --- | --- | --- | --- |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |
|  |  | [DD/MM/YYY] |  |

### Contents

Molecular Testing for Malaria: Overview of Standards	3

Signature Page	4

Version History	5

Contents	6

1.	Scope	8

2.	Abbreviations	8

3.	Personnel qualifications	8

3.1.	Medical fitness	8

3.2.	Education and training	8

4.	Procedure	9

4.1.	Principle	9

4.2.	Samples	9

4.3.	Required Reagents, Materials and Equipment	9

4.4.	Supplies	9

4.5.	Equipment	9

4.6.	Master Mix Reagents	10

4.7.	Procedural steps	10

5.	Quality Control	15

5.1.	Negative Controls	15

5.2.	Positive Controls	15

6.	Procedure limitations	16

7.	Interpretation and Reporting of Results	16

8.	Safety Precautions	16

9.	Bibliography	18

10.	Appendices	20

Appendix A – PCR Reagent Concentration and Master Mix Tools	21

Appendix B – Bench Top Reference	23

Appendix C – Master Mix Worksheet	27

Appendix D –PCR result Worksheet	28

Appendix E – Recommended msp1 Primer Sets	29

# Scope

This SOP describes the method used to amplify the msp1 gene from extracted Plasmodium falciparum DNA. The method applies to day zero and post-therapy human blood samples used for the determination of recrudescence versus reinfection.

# Abbreviations

dNTP	Deoxynucleotidetriphosphate

DBS	Dried Blood Spot

MgCl2	Magnesium Chloride

MSP1	Merozoite Surface Protein 1

M	micromolar

ml	milliliter

mM	milimolar

msp1	The gene encoding MSP1

nPCR	Nested PCR

Pf	Plasmodium falciparum

pPCR	Primary PCR

PCR	Polymerase chain reaction

s	Second

SOP	Standard Operating Procedure

L	microlitre

# Personnel qualifications

## Medical fitness

Occupational health programs should be in place to monitor/address staff vaccinations and deal with exposures to potentially infected materials.

## Education and training

Training must be given on the following topics:

- Wearing and use of personal protective equipment and clothing;

- Handling of potentially infectious materials;

- Prevention of incidents and steps to be taken by workers in case incidents (including biological, chemical, electrical and fire hazards) occur;

- Procedures;

- Waste management;

- Impact of results for patient management and research.

Training must be provided:

- When a new staff member takes up post;

- Annually;

- When there is a change in conditions or best practices.

# Procedure

## Principle

The most widely used genetic markers for malaria genotyping are the antigen genes msp1, msp2 and glurp.  These markers have been shown to have adequate discriminatory power for recrudescence versus reinfection testing (World Health Organization, 2007).

This protocol uses two stages of PCR to increase sensitivity and specificity for msp1. The first stage is a primary PCR (pPCR) in which the msp1 gene is amplified. The second stage is a nested PCR (nPCR) in which strain-specific primers are used to amplify from the pPCR template.

There is one set of primers for the pPCR and 3 sets of primers for the nPCR. The primers for the nPCR amplify 3 common alleles from the following Pf strains: K1, MAD20 and RO33. Consequently, each pair of samples (day 0 and post-treatment) will be subjected to pPCR with each set of primers.

## Samples

Parasite DNA samples are obtained using the DBS DNA extraction protocols of SOP-3 or SOP-04. Extracted DNA is ready for use in PCR.

## Required Reagents, Materials and Equipment

## Supplies

1.5 ml microcentrifuge tubes and racks

0.2 ml PCR tubes (with cap) and racks

1-20 l, 100-200 l and 1000 l single channel automatic pipettes

Filter pipette tips for the above pipettes

Mineral oil (sterile) (if using thermocycler without heated lid)

Fine tip marker pens

Disposable gloves

## Equipment

Microcentrifuge

Vortex

Thermocycler

## Master Mix Reagents

Nuclease-free water – maintain 500-1000 l aliquots at room temperature.

PCR buffer (MgCl2-free) (composition varies) – maintain 100 l aliquots at -20ºC.

MgCl2 (concentration varies) – maintain 50 l aliquots at -20ºC.

dNTP (concentration varies) – maintain 50 l aliquots at -20ºC.

Taq DNA Polymerase – maintain 50 l aliquots at -20ºC.

Primers for msp1 pPCR and nPCR K1, MAD20 and RO33 (see appendix E for recommended primer sets) – make stock solutions (e.g. 5, 10 or 25 mM) and maintain in 50 l aliquots at -20ºC.

NB. See PCR Reagent Concentration and Master Mix Tools in Appendix A.

## Procedural steps

Important points to remember: - Never used expired reagents as it is not possible to assure accuracy of results. - There are two stages of PCR, the primary PCR (pPCR) and the nested PCR (nPCR). - Preparation of reagents must be performed in clean area, separate from amplification. - Employ stringent cleaning procedures to limit contamination. - Keep a separate set of single channel pipettes for reagent preparation. These are not for any other use! - Aliquot large quantities of PCR reagents and store appropriately. - Always include a DNA extraction (N0), positive (P1 and P2), and negative (N1 and N2) controls with each batch (see section 5 for more details). - Refer to Appendices for information on primers and dilutions.

- Ensure the thermocycler is available for the appropriate amount of time.

- Print out a master mix worksheet (Appendix C). Either print out a PCR worksheet (Appendix D) and record the ID number of each sample to be tested on a separate numbered line OR use the printed worksheet from the PCR of msp1 or glurp. Results for each batch should be kept on the same worksheet. NB. Keep day 0 and post-therapy samples from each patient in numerical order.

- Record the DNA extraction (N0), positive (P1), positive (P2) and negative (N1) controls on the worksheet (if this hasn’t been done previously). NB. All controls are essential to assure quality testing (see section 5 for details).

Stage 1: pPCR - Steps Performed in the Master Mix / Clean Room - Wash hands with soap and water and put on gloves. - Thoroughly clean work area and all equipment (including pipettes) with water. - Gather all required supplies and remove master mix reagents from the freezer to thaw at room temperature. NB. Only remove Taq Polymerase from freezer when required. Taq Polymerase is the last reagent to be added to the master mix. - After thawing, vortex* (5 s) and centrifuge (5 s) reagents. *Taq polymerase should be mixed by gently flicking the side of the tube a few times, it should not be vortexed. - Label sufficient number of 0.2 ml microcentrifuge tubes according to the sample order on the worksheet. NB. If possible, include sample ID number as well as worksheet position. - Label a 0.2 ml tube for each of the following controls: N0, N1 and P1. - Prepare Master Mix in a 1.5 ml microcentrifuge tube (labeled pPCR MM) according to the worksheet below (also see Appendices C and D for template worksheets). Vortex, add Taq, flick and briefly centrifuge. Reagent (pPCR); Final Conc.; Vol. x1 sample; Vol. x N samples; Check / tick when added. 1.; Nuclease-free water; Up to 18 l. 2.; PCR Buffer; 1 x. 3.; MgCl2; 2 mM. 4.; dNTP; 125 µM. 5.; Primer - forward; 125 nM. 6.; Primer - reverse; 125 nM. 7.; Taq Polymerase; 0.4 U. Total; 18 l. NB. See Appendix A for information on master mix calculations. Complete and store worksheet for each batch of PCR reactions. - Return all stock reagents to the freezer. - Mix by flicking the side of the tube a few times and briefly centrifuge (5 s). - Aliquot 18 l master mix into each labeled 0.2 ml tube (for samples and controls). - Ensure work area is cleaned and all pipettes are set to the maximum volume before moving on to next step. NB. When using a thermocycler without a heated lid, add 25 l mineral oil to each PCR tube. Keep PCR tube lids closed in between addition of reagents and when transporting samples between preparation areas.

Stage 1: pPCR - Steps Performed in the DNA Room NB. PCR facilities may vary. In some cases a separate DNA room may not be available. It is recommended that master mix preparation and addition of template DNA be kept as separate as possible (i.e. separate rooms or at least separate areas and pipettes). In ALL cases, amplified products should be kept completely separate from master mix preparation areas (always in a separate room with dedicated equipment and supplies). - Wash hands with soap and water and put on gloves. - Using a new pipette tip each time, add 2 l sample DNA to the appropriately-labeled tube. NB. Take great care not to create aerosols when pipetting. Keep PCR tube lids closed in between addition of reagents and when transporting samples between preparation areas. - Add control samples: 2 l nuclease-free water to N1, 2 l of the positive control (see section 5) to P1 and 2 l of the DNA extraction negative control to N0. - Perform the pPCR in the thermocycler according to the following program: Step no.; Cycle; Temperature (ºC); Time (min); No. of cycles. 1; Initial Denaturation; 95; 5; 1. 2; Denaturation; 94; 1; 25. 3; Annealing; 58; 2. 4; Extension; 72; 2. 5; Final extension; 72; 5; 1. NB. Reactions should be removed from the thermocycler and stored at +4 ºC until ready for use in the nPCR (second stage). For periods longer than 48 hrs, samples should be stored at -20 ºC. Turn the thermocycler off after use.

Stage 2: nPCR - Steps Performed in the Master Mix / Clean Room - Set up one master mix at a time for each pPCR primer pair (K1, MAD20 and RO33). - Wash hands with soap and water and put on gloves. - Thoroughly clean work area and all equipment (including pipettes) with water. - Gather all required supplies and remove master mix reagents from the freezer to thaw at room temperature. NB. Only remove Taq Polymerase from freezer when required. Taq Polymerase is the last reagent to be added to the master mix. - After thawing, vortex* (5 s) and centrifuge (5 s) each. *Taq polymerase should be mixed by gently flicking the side of the tube a few times, it should not be vortexed. - Label sufficient number of 0.2 ml microcentrifuge tubes according to the sample order on the worksheet. Indicate “nPCR” on each tube. NB. If possible, include sample ID number as well as worksheet position. - Label the control tubes: N0 nPCR (for the DNA extraction negative control), N2 (for N1 template) and P2 (for P1 template). - Prepare Master Mix in a 1.5 ml microcentrifuge tube (labeled nPCR MM) according to the worksheet below (See Appendices C and D for full range template worksheets). Vortex, add Taq, flick and briefly centrifuge. Reagent (nPCR); Final Conc.; Vol. x1 sample; Vol. x N samples; Check / tick when added. 1.; Nuclease-free water; Up to 18 l. 2.; PCR Buffer; 1 x. 3.; MgCl2; 2 mM. 4.; dNTP (each); 125 µM. 5.; Primer - forward; 250 nM. 6.; Primer - reverse; 250 nM. 7.; Taq Polymerase; 0.4 U. Total; 18 l. NB. See Appendix A for information on master mix calculations and stock solutions. Complete and store worksheet for each batch of PCR reactions. - Return all stock reagents to the freezer. - Mix by flicking the side of the tube a few times and briefly centrifuge (5 s).

- Aliquot 18 l master mix into each labeled 0.2 ml tube (for samples and controls). NB. Add 25 l mineral oil to each PCR tube. NB. Keep PCR tube lids closed in between addition of reagents and when transporting samples between preparation areas - Ensure work area is cleaned and all pipettes are set to the maximum volume before moving on to next step.

Stage 2: nPCR - Steps Performed in the DNA/Amplification Room NB. Master mix for nPCR is taken to the DNA/Amplification room for addition of template DNA. Amplified DNA should never be introduced into clean rooms. - Wash hands with soap and water and put on gloves. - Using a new pipette tip each time, add 2 l template DNA (from the pPCR) to the appropriately-labeled tube. NB. Take great care not to create aerosols when pipetting. Keep PCR tube lids closed in between addition of reagents and when transporting samples between preparation areas. - Add control samples – 2 l N1 into N2, P1 into P2 and N0 into N0 nPCR. - Perform the nPCR in the thermocycler according to the following program: Step no.; Cycle; Temperature (ºC); Time (min); No. of cycles. 1; Initial Denaturation; 95; 5; 1. 2; Denaturation; 94; 1; 30. 3; Annealing; 61; 2. 4; Extension; 72; 2. 5; Final extension; 72; 5; 1. NB. Reactions should be removed from the thermocycler and stored at +4 ºC until ready for use in the nPCR (second stage). For periods longer than 48 hrs, samples should be stored at -20 ºC. Turn the thermocycler off after use.

- Repeat steps 19-33 using primers for each allele (K1, MAD20 and RO33).

- The samples are ready for gel electrophoresis according to SOP-08.

# Quality Control

## Negative Controls

The negative controls employed in this SOP are as follows:

| Control ID | Description | Purpose | Corrective Action (in event of failure) |
| --- | --- | --- | --- |
| N0 | Plain filter paper negative control made during DBS DNA extraction (SOP-03 or SOP-04). | To check DNA extraction process for cross-contamination of samples. | - Clean scissors/punch - Clean pipettes - Use new filter tip with each sample - Discard water and use fresh aliquot |
| N1 | A negative control for pPCR in which water is added instead of template DNA. | To check pPCR process for cross-contamination of samples. Keep refrigerated and only use if N2 control fails. | - Clean pipettes - Use new filter tip with each sample - Discard water, buffer, dNTP, primer and remake. |
| N2 | A negative control for nPCR in which N1 is added instead of template DNA. | To check nPCR process for cross-contamination of samples. | - Clean pipettes - Use new filter tip with each sample - Discard water, buffer, dNTP, primer and remake. |

## Positive Controls

The positive controls employed in this SOP are as follows:

| P1 | A positive control for pPCR in which control DNA is added instead of template DNA. | To check that all reagents and pPCR conditions work. | - Discard water, buffer, dNTP, primer and remake. - Use new Taq Polymerase. |
| --- | --- | --- | --- |
| P2 | A positive control for nPCR in which control P1 is added instead of template DNA. | To check that all reagents and nPCR conditions work. | - Discard water, buffer, dNTP, primer & remake. - Use new Taq Polymerase. |

Positive control stock

Positive control stock solutions can be made using purified Pf DNA. The table below displays strains that can be used in each PCR.  marks indicate where strains template DNA is suitable for PCR.

| Template DNA | Locus / Strain |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | msp1 |  |  | msp2 |  | glurp |
|  | K1 | MAD20 | RO33 | FC27 | 3D7/IC | N/A |
| 3D7 |  |  |  |  |  |  |
| 7G8 |  |  |  |  |  |  |
| D6 |  |  |  |  |  |  |
| D10 |  |  |  |  |  |  |
| Dd2 |  |  |  |  |  |  |
| FC27 |  |  |  |  |  |  |
| HB3 |  |  |  |  |  |  |
| K1 |  |  |  |  |  |  |
| MAD20 |  |  |  |  |  |  |
| RO33 |  |  |  |  |  |  |
| W2 |  |  |  |  |  |  |

DNA from each strain should be kept in separate 10 l aliquots of 1 ng ul-1 stock. To make 100 l of 25 pg ul-1 positive control working solution, add 2.5 l of the 1 ng ul-1 stock to 97.5 l water.  All control stock and working solutions should be kept at -20 ºC. See Appendix A for details on making stock solutions.

# Procedure limitations

- The quality of the DBS sample is of critical importance for every step of recrudescence versus reinfection testing.

- Poor quality samples and/or insufficient DNA extraction may result in a lack of PCR products and/or spurious bands.

- The appropriate quality-assured storage of reagents and samples is also critical to ensure the highest quality PCR results.

# Interpretation and Reporting of Results

Results are reported on the worksheet in Appendix D after gel electrophoresis of PCR products (see SOP-08).

# Safety Precautions

Although extracted DNA is a low biohazard risk specimen, always practice universal precautions (treat all patient specimens as potentially infectious material:

- Wear good quality, single-use, disposable medical examination gloves.

- Wear a laboratory coat or gown.

- Wash hands after removal of gloves.

Universal precautions always apply in the laboratory environment. In addition, laboratory staff should:

- Wear sturdy, closed-toe shoes;

- Put on protective eye wear for high-risk activities such as working with UV light;

- Refrain from drinking, eating, smoking and storing food anywhere in the laboratory; and

- Remove gloves and wash hands before touching clean areas such as door handles, lift/elevator buttons and telephones.

# Bibliography

Calderaro, A., Piccolo, G., Perandin, F., Gorrini, C., Peruzzi, S., Zuelli, C., et al. (2007). Genetic polymorphisms influence Plasmodium ovale PCR detection accuracy. Journal of Clinical Microbiology, 45(5), 1624-1627.

Cattamanchi, A. A., Kyabayinze, D., Hubbard, A., & Rosenthal, P. J. (2003). Distinguishing recrudescence from reinfection in a longitudinal antimalarial druf efficacy study: comparison of results based on genotyping of msp-1, msp-2, and glurp. American Journal of Tropical Medicine, 68(2), 133-139.

Centers for Disease Control and Prevention. (2006). Blood collection - finger prick. Retrieved August 17, 2010, from CDC HIV Training: http://wwwn.cdc.gov/dls/ila/hivtraining/trainersguide/pdf/presentations/Module8Presentation.pdf

Clinical and Laboratory Standards Institute. (2007). Blood collection on filter paper for newborn screening programs; approved standard - fifth edition. CLSI document LA4-A5, 27(20).

Falk, N., Maire, N., Sama, W., Owusu-Agyei, S., Smith, T., & Beck, H.-P. a. (2006). Comparison of PCR-RFLP and Genescan-based genotyping for analyzing infection dynamics of Plasmodium falciparum. American Journal of Tropical Medicine., 74(6), 944-950.

Färnert, A., Arez, A. P., Babiker, H. A., Beck, H. P., Benito, A., Björkman, A., et al. (n.d.).

Färnert, A., Arez, A. P., Babiker, H. A., Beck, H. P., Benito, A., Björkman, A., et al. (n.d.). Genotyping of Plasmodium falciparum infections by PCR: a comparative multicentre study. Trans R Soc Trop Med Hyg., 95.

Kyes, S., Craig, A. G., & Marsh, K. a. (1993). Plasmodium falciparum: a method for the amplification of S antigens and its application to laboratory and field samples. Experimental Parasitology, 71, 473-483.

Nakeesathit, S., Pagomrat, W., Tanomsing, N., & Hanchana, S. a. (2001). DNA Extraction. Bangkok: Mahidol Oxford Tropical Medicine Research Unit.

Plowe, C. V., Djimde, A., Bouare, M., & Doumbo, O. a. (1995). Pyrimethamine and proguanil resistance-conferring mutations in Plasmodium falciparum dihydrofolate reductase: polymerase chain reaction methods for surveillance in Africa. American Journal of Tropical Medicine and Hygiene., 52, 565-568.

Program for Appropriate Technology in Health (PATH). (2005). RBP-EIA: collecting, processing, and handling venous, capillary, and blood spot samples. Seattle: PATH.

QIAGEN. (2007). QIAamp DNA mini and blood mini handbook Second edition. QIAGEN.

Rubio, J. M., J., R. P., L., B. M., Garcia, M., M., M., & I., E. M. (1999). Semi-nested. multiplex polymerase chain reaction for detection of human malaria parasites and evidence of Plasmodium vivax infection in Equatorial Guinea. American Journal of Tropical Medicine and Hygiene., 60, 183-187.

U.S. Department of Health and Human Services. (2007). Biosafety in microbiological and biomedical laboratories. 5th. Washington, District of Columbia, USA: U. S. Government Printing Office.

U.S. Department of Health and Human Services. (2007). Biosafety in microbiological and biomedical laboratories (5th ed.). Washington, District of Columbia, USA: U. S. Government Printing Office.

Watcharee, P., Naowarat, T., Supatchara, N., & Sarun, H. a. (2009). Gel Electrophoresis. Bangkok: MORU.

World Health Organization. (2007). Methods and techniques for clinical trials on antimalarial drug efficacy: genotyping to identify parasite populations. Amsterdam: WHO.

World Health Organization. (2007). Recommended Genotyping Procedures (RGPs) to identify parasite populations. Amsterdam: Medicines for Malaria Venture and World Health Organization.

World Health Organization. (2008). Consultation on Technical and Operational Recommendations for Clinical Laboratory Testing Harmonization and Standardization. Geneva: World Health Organization.

Worldwide Antimalarial Resistance Network. (2010). Filter paper preparation v1.0 (SOP ID: MOL03/CLIN06). Worldwide Antimalarial Resistance Network (WWARN).

Zwetyenga, J., Rogier, C., Tall, A., Fontenille, D., Snounou, G., & Trape, J.-F. a.-P. (1998). No influence of age on infectious complexity and allelic distribution in Plasmodium falciparum infections in Ndiop, a Senegalese village with seasonal, mesoendemic malaria. American Journal of Tropical Medicine and Hygiene., 59(5), 726-735.

# Appendices

### Appendix A – PCR Reagent Concentration and Master Mix Tools

- Making stock solutions

To dilute a stock solution into a working solution remember:

| What you want | X Volume |
| --- | --- |
| What you have |  |

Before calculating, ensure all units are the same (multiple or divide by appropriate factor):

|  | pg l-1or pM | ng l-1or nM | g l-1 or M | mg l-1 or mM |
| --- | --- | --- | --- | --- |
| pg l-1 or pM | 1 | 103 | 106 | 109 |
| ng l-1 or nM | 103 | 1 | 103 | 106 |
| g l-1 or M | 106 | 103 | 1 | 103 |
| mg l-1 or mM | 109 | 106 | 103 | 1 |

Example:

I have a stock solution of 1 ng l-1 and want to make a 50 l working solution of 25 pg l-1.

Step 1: Change the 1 ng l-1 to pg l-1:

From the chart, the difference is a factor 103 therefore,

1 ng l-1 = 1 x 103 pg l-1 = 1000 pg l-1

Step 2: Plug the numbers into the equation:

| 25 pg l-1 (what you want) | 50 l (volume you want) | = | 1.25 l |
| --- | --- | --- | --- |
| 1000 pg l-1 (what you have) |  |  |  |

Therefore, take 1.25 l stock and add 48.75 l to make 50 l of a 25 pg l-1 working solution.

- PCR Tools

There are a range of tools available online to assist in calculating the concentrations and required volume of reagents in a master mix based on the number of samples. A sample of tools can be found at:

- Sigma-Aldrich PCR MasterMix Calculator: http://www.sigmaaldrich.com/life-science/molecular-biology/pcr/learning-center/pcr-tools.html

- Finnzymes OD Conversion: http://www.finnzymes.com/java_applets/od_conversion.html

- Promega DNA and Temperature Conversions and Conversion Factors for Molecular Biology: http://www.promega.com/biomath/Default.htm

### Appendix B – Bench Top Reference

Molecular Testing for Malaria (MTM) Bench Top Reference

Document ID: BTR-05 msp1 – pPCR Prep in Clean Room

- Print out a PCR worksheet and record the ID number of each sample to be tested on a separate numbered line. Keep pairs in sequence.

- Record the DNA extraction (N0), positive (P1), positive (P2) and negative (N1) controls on the worksheet.

Stage 1: pPCR - Steps Performed in the Master Mix / Clean Room - Wash hands with soap and water and put on gloves. - Thoroughly clean work area and all equipment (including pipettes) with water. - Gather all required supplies and remove master mix reagents from the freezer to thaw at room temperature. - After thawing, vortex* (5 s) and centrifuge (5 s) reagents for a few seconds each. *Do not vortex Taq polymerase! - Label sufficient number of 0.2 ml microcentrifuge tubes according to the sample order on the worksheet. - Label a 0.2 ml tube for each of the following controls: N0, N1 and P1. - Prepare Master Mix in a 1.5 ml microcentrifuge tube (labeled pPCR MM) according to the worksheet below: Reagent (pPCR); Final Conc.; Vol. x1 sample; Vol. x N samples; Check / tick when added. 1.; Nuclease-free water; Up to 18 l. 2.; PCR Buffer; 1 x. 3.; MgCl2; 2 mM. 4.; dNTP; 125 µM. 5.; Primer - forward; 125 nM. 6.; Primer - reverse; 125 nM. 7.; Taq Polymerase; 0.4 U. Total; 18 l. - Return all stock reagents to the freezer. - Mix by flicking the tube a few times and briefly centrifuge (5 s). - Aliquot 18 l master mix into each labeled 0.2 ml tube (for samples and controls). - Add 25 l mineral oil to each PCR tube. - Ensure work area is cleaned and all pipettes are set to the maximum volume before moving on to next step.

Molecular Testing for Malaria (MTM) Bench Top Reference

Document ID: BTR-05 msp1 – pPCR Prep in DNA Room

Stage 1: pPCR - Steps Performed in the DNA Room NB. PCR facilities may vary. In some cases a separate DNA room may not be available. It is recommended that master mix preparation and addition of template DNA be kept as separate as possible (i.e. separate rooms or at least separate areas and pipettes). In ALL cases, amplified products should be kept completely separate from master mix preparation areas (always in a separate room with dedicated equipment and supplies). - Wash hands with soap and water and put on gloves. - Using a new pipette tip each time, add 2 l sample DNA to the appropriately-labeled tube. NB. Take great care not to create aerosols when pipetting. Keep PCR tube lids closed in between addition of reagents and when transporting samples between preparation areas. - Add control samples – 2 l nuclease-free water to N1, 2 l of the positive control (see section 5) to P1 and 2 l of the DNA extraction negative control to N0. - Perform the pPCR in the thermocycler according to the following program: Step no.; Cycle; Temperature (ºC); Time (min); No. of cycles. 1; Initial Denaturation; 95; 5; 1. 2; Denaturation; 94; 1; 25. 3; Annealing; 58; 2. 4; Extension; 72; 2. 5; Final extension; 72; 5; 1. NB. Reactions should be removed from the thermocycler and stored at +4 ºC until ready for use in the nPCR (second stage). For periods longer than 48 hrs, samples should be stored at -20 ºC.Turn the thermocycler off after use.

Molecular Testing for Malaria (MTM) Bench Top Reference

Document ID: BTR-05 msp1 – nPCR Prep in Clean Room

Stage 2: nPCR - Steps Performed in the Master Mix / Clean Room - Wash hands with soap and water and put on gloves. - Thoroughly clean work area and all equipment (including pipettes) with water. - Gather all required supplies and remove master mix reagents from the freezer to thaw at room temperature. NB. Only remove Taq Polymerase from freezer when required. Taq Polymerase is the last reagent to be added to the master mix. - After thawing, vortex* (5 s) and centrifuge (5 s) reagents for a few seconds each. *Do not vortex Taq polymerase! - Label sufficient number of 0.2 ml microcentrifuge tubes according to the sample order on the worksheet. Indicate “nPCR” and the allele (e.g. K1) on each tube. NB. If possible, include a second sample ID number as well as the worksheet position. - Label the control tubes: N0 nPCR (for the DNA extraction negative control), N2 (for N1 template) and P2 (for P1 template). - Prepare Master Mix in a 1.5 ml microcentrifuge tube (labeled nPCR MM) according to the worksheet below: Reagent (nPCR); Final Conc.; Vol. x1 sample; Vol. x N samples; Check / tick when added. 1.; Nuclease-free water; Up to 18 l. 2.; PCR Buffer; 1 x. 3.; MgCl2; 2 mM. 4.; dNTP (each); 125 µM. 5.; Primer - forward; 250 nM. 6.; Primer - reverse; 250 nM. 7.; Taq Polymerase; 0.4 U. Total; 18 l. - Return all stock reagents to the freezer. - Mix by flicking the tube a few times and briefly centrifuge (5 s). - Aliquot 18 l master mix into each labeled 0.2 ml tube (for samples and controls). - If required (for thermocyclers without heated lids), add 25 l mineral oil to each PCR tube. Ensure work area is cleaned and all pipettes are set to the maximum volume before moving on to next step.

Molecular Testing for Malaria (MTM) Bench Top Reference

Document ID: BTR-05 msp1 – nPCR Prep in DNA Room

Stage 2: nPCR - Steps Performed in the DNA/Amplification Room NB. Master mix for nPCR is taken to the DNA/Amplification room for addition of template DNA. Amplified DNA should never be introduced into clean rooms. - Wash hands with soap and water and put on gloves. - Using a new pipette tip each time, add 2 l template DNA (from the pPCR) to the appropriately-labeled tube. NB. Take great care not to create aerosols when pipetting. Keep PCR tube lids closed in between addition of reagents and when transporting samples between preparation areas. - Add control samples – 2 l N1 into N2, P1 into P2 and N0 into N0 nPCR. - Perform the nPCR in the thermocycler according to the following program: Step no.; Cycle; Temperature (ºC); Time (min); No. of cycles. 1; Initial Denaturation; 95; 5; 1. 2; Denaturation; 94; 1; 30. 3; Annealing; 61; 2. 4; Extension; 72; 2. 5; Final extension; 72; 5; 1. NB. Reactions should be removed from the thermocycler and stored at +4 ºC until ready for use in the next stage. For periods longer than 48 hrs, samples should be stored at -20 ºC.Turn the thermocycler off after use.

- The samples are ready for gel electrophoresis according to SOP-08.

### Appendix C – Master Mix Worksheet

| Date: |  |  |  |  | Initials: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| msp1 pPCR |  |  |  |  |  |  |  |  |  |  |  |
| Reagent |  |  | Final Conc. | Vol. x1 sample |  | Vol. x N samples |  | Check / tick when added |  |  |  |
| 1. |  | Nuclease-free water | Up to 18 l |  |  |  |  |  |  |  |  |
| 2. |  | PCR Buffer | 1 x |  |  |  |  |  |  |  |  |
| 3. |  | MgCl2 | 2 mM |  |  |  |  |  |  |  |  |
| 4. |  | dNTP (each) | 125 µM |  |  |  |  |  |  |  |  |
| 5. |  | Primer - forward | 125 nM |  |  |  |  |  |  |  |  |
| 6. |  | Primer - reverse | 125 nM |  |  |  |  |  |  |  |  |
| 7. |  | Taq Polymerase | 0.4 U |  |  |  |  |  |  |  |  |
|  |  | Total |  | 18 l |  |  |  |  |  |  |  |
| msp1 nPCR |  |  |  |  |  |  |  |  |  |  |  |
| Reagent |  |  | Final Conc. | Vol. x1 sample |  | Vol. x N samples | Check / tick when added |  |  |  |  |
|  |  |  |  |  |  |  | K1 |  |  | MAD20 | RO33 |
| 1. | Nuclease-free water |  | Up to 18 l |  |  |  |  |  |  |  |  |
| 2. | PCR Buffer |  | 1 x |  |  |  |  |  |  |  |  |
| 3. | MgCl2 |  | 2 mM |  |  |  |  |  |  |  |  |
| 4. | dNTP (each) |  | 125 µM |  |  |  |  |  |  |  |  |
| 5. | Primer - forward |  | 250 nM |  |  |  |  |  |  |  |  |
| 6. | Primer - reverse |  | 250 nM |  |  |  |  |  |  |  |  |
| 7. | Taq Polymerase |  | 0.4 U |  |  |  |  |  |  |  |  |
|  | Total |  |  | 18 l |  |  |  |  |  |  |  |

### Appendix D –PCR result Worksheet

| Date |  | Staff Initials |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Number | Sample ID | PCR Product Sizes (bp) for control, indicate band sizes or leave blank |  |  |  |  |  |  |  |  |  |  | Comments & Results |
|  |  | msp1 |  |  |  |  | msp2 |  |  |  | glurp |  |  |
|  |  | K1 |  | MAD20 | RO33 |  | FC27 |  | 3D7/IC |  |  |  |  |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 11 | N0 |  |  |  |  |  |  |  |  |  |  |  |  |
| 12 | N1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 13 | P1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 14 | N2 |  |  |  |  |  |  |  |  |  |  |  |  |
| 15 | P2 |  |  |  |  |  |  |  |  |  |  |  |  |

### Appendix E – Recommended msp1 Primer Sets

| User / Source | PCR | Locus | Allele | Primer |  | Sequence (primer differences are shown in bold, underlined font) | Reference |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Forward | Reverse |  |  |
| WHO, MORU, IPC and UMB | pPCR | msp1 | N/A |  |  | CTAGAAGCTTTAGAAGATGCAGTATTG | WHO primers are taken from (World Health Organization, 2007). |
|  |  |  |  |  |  | CTTAAATAGTATTCTAATTCAAGTGGATCA |  |
|  | nPCR |  | K1 |  |  | AAATGAAGAAGAAATTACTACAAAAGGTGC |  |
|  |  |  |  |  |  | GCTTGCATCAGCTGGAGGGCTTGCACCAGA |  |
|  |  |  | MAD20 |  |  | AAATGAAGGAACAAGTGGAACAGCTGTTAC |  |
|  |  |  |  |  |  | ATCTGAAGGATTTGTACGTCTTGAATTACC |  |
|  |  |  | RO33 |  |  | TAAAGGATGGAGCAAATACTCAAGTTGTTG |  |
|  |  |  |  |  |  | CAAGTAATTTTGAACTCTATGTTTTAAATCAGCGTA CATCTGAAGGATTTGCAGCACCTGGAGATC |  |

It is recommended to aliquot out and store all new reagents. The required volume of each aliquot depends upon workload and subsequent PCR batch size. These are general recommendations.

Make sufficient Master Mix for 2 extra reactions to account for error.

Keep on ice until ready to use.

Make sufficient Master Mix for 2 extra reactions to account for error.

Keep on ice until ready to use.